## Slide 1: Software for Embedded Systems
Prof. Saurabh Gandhi (EEE department)

## Slide 2: Interrupts
Hardware events capture CPU attention
Normal execution is temporarily suspended
A function (“handler”) is called to handle the event
Normal execution resumes after event is handled
Hardware decides when the software reacts
More responsive!

## Slide 3: What does it mean to ‘handle an event’?
Examples:
A keyboard key press
New temperature measurement available in a weather logger
New temperature measurement available in the infusion pump
Handling interrupts correctly is critical – do the bare minimal for survival (and design survival accordingly)

## Slide 4: Handling multiple interrupt sources
What if an interrupt comes while handling another interrupt?
Infusion pump:
  temperature sensor interrupt comes while handling user input
  User input interrupt comes while handling temperature sensor
  Temperature sensor vs bubble sensor
Disabling interrupts
Priorities
Hybrid model: poll slowly changing signals, and use interrupts for critical events

## Slide 5: Modeling embedded systems – UML
What do we mean by ‘modeling’ an ES?
What is UML (Unified Modeling Language)?
Why UML for embedded systems?
System context, boundaries and use cases

## Slide 6: What does it mean to ‘model’ a system?
Models are an abstraction of reality
Partial, imperfect descriptions
What is the use of partial and imperfect descriptions?
Model of the Delhi metro
Which is most accurate?
Which of these is most useful?
Use dependent utility

## Slide 7: What is a model?
A model is an abstraction of a physical system. It highlights the essential characteristics while hiding unnecessary implementation details.
The "Blueprint" Analogy:
  You wouldn't build a skyscraper without a structural model
  You shouldn't build a safety-critical embedded system without a behavioral model
Multiple views of a model
Structural: what the system is made of
Behavioral: How the system reacts to stimuli
Environmental: How the system sits in the real world (what it senses, what it affects, who uses it etc.)
Models help capture assumptions early on!

## Slide 8: Why model embedded systems?
Complexity management: Modern embedded systems (e.g., an Autonomous Driving module) have millions of lines of code. Humans cannot hold the entire logic in their heads at once
Early error detection: Finding a logic flaw in a UML diagram costs cents; finding it during "Hardware-in-the-Loop" testing costs thousands; finding it after a product recall costs millions
Communication bridge between H/W and S/W engineers; designers and stakeholders
Basis for simulation: Models can be mathematically verified or simulated to prove certain claims (for example, related to safety)

## Slide 9: Without models
System boundaries drift (who handles a sensor fault? Not defined, and no one does it…)
Responsibilities are unclear (both the control task and interrupt try to stop a motor, causing race conditions)
Assumptions remain implicit (Siemens PLC example)
Interfaces are poorly defined (sensor driver reports raw values and control module assumes another unit)
Event handling is incomplete (over-temperature event is handled, but recovery is undefined)
Fault cases are forgotten (normal infusion works perfectly, but behavior after watchdog timer reset is never defined)
Verification becomes ad hoc
Changes break unrelated behavior (adding a new power mode disables alarms because wake-up assumptions were not modeled)

## Slide 10: The Unified Modeling Language (UML)
A standardized, graphical language for visualizing system architecture (structure, behavior and environment)
Specifically for embedded systems: SysML (a subset of UML that focuses on real-time constraints and hardware interfaces)
The three pillars of UML:
  Things (players involved in the system)
  Relationships (associations / dependencies between them)
  Diagrams (snapshots of the model from different angles)

## Slide 11: System context: The first modeling question
What is the system, and what is not the system?
System boundary defines what is inside and what is outside your system
Whatever falls outside the boundary is an assumption
Whatever falls inside is a specification to be designed and verified

## Slide 12: System boundary: Infusion pump example
Nurse? Motor driver?
Stepper motor?
  Why outside: software responsibility is to send pulses. If motor jams, not software responsibility
  Why inside: They system’s purpose is not to calculate motor angle and send it; purpose is to deliver right amount of drug
No correct answer, but important to agree on an answer! Otherwise…
  Control logic will assume motor is behaving as expected
  Motor jams are not detected as faults
  Software not designed to handle motor faults
What about the temperature sensor?

## Slide 13: Actors – who or what interacts with the system?
Anything external that interacts with the system
  Human users
  External ‘hardware’ (sensors, objects etc.)
    …air bubble?
  External systems (cloud server)
Actors are determined by the system boundary (for example if sensors are inside the system, they are not actors)

## Slide 14: External interfaces
How actors influence the system
Physical interfaces: GPIO pins, communication protocols (SPI, I2C, UART etc.)
Logical interfaces: APIs, protocols (MQTT, HTTP)
The context diagram: high level view showing the system at the center and all actors connected via interfaces

## Slide 15: Use Cases: meaningful interactions between actors and system
Use case describes a goal-oriented interaction between an actor and the system
Describe behavior, not interface design!
Behavioral contract: what the system must do in response to a stimulus (not about how it does it)
Example:
  User clicks ‘stop’ and screen turns red (WRONG)
  System shall transition to a safe state within 10 ms of receiving the ‘stop’ signal from user (CORRECT)
Use case defines
  Actor
  Trigger
  Main success scenario
  Alternatives / exception scenarios (fault recovery behavior)

## Slide 16: Use cases – infusion pump example
Example use case: deliver drug at programmed infusion rate safely
Who are the actors for this use case?
  Nurse (primary), power source, sensors, motor…
What is the trigger?
  Nurse programs an infusion rate and starts infusion
Main success scenario:
  System accepts programmed infusion rate
  System starts pump motor
  System continuously regulates flow
  System monitors safety during infusion (detects occlusions or over-temperature)
  System completes infusion and stops motor
Alternate scenarios:
  Over temperature detected: system stops motor and raises alarm
  Occlusion detected…
  Sensor failure detected…

## Slide 17: Use cases and requirements in the V-Model
One use case  many requirements
One requirement  supports multiple use cases
Sit near top-left of V-Model
Help define system tests and acceptance criteria

## Slide 18: Summary
UML is a reasoning language
System boundaries define responsibility
Actors are external to the system
Use cases describe behavioral contracts, not UI flows
Correct context modeling prevents downstream errors

## Slide 19: Behavioral modeling with UML statecharts
Embedded systems are event-driven
Statecharts allow reasoning about system behavior as a whole
Ensures for every state and input there is a unique, well-defined outcome
Prove system cannot enter into a deadly / undefined state

## Slide 20: States, events and transitions
State: represents a specific, finite condition or mode in an object's lifecycle, where it exhibits consistent behavior in response to events
Event: a significant occurrence, stimulus, or message that triggers a transition from one state to another
  Interrupt driven (triggered by external actors)
  Time-triggered
Transition: change from one state to another driven by an event

## Slide 21: States and events in the infusion pump
States
  Idle
  Infusing
  Fault
Events
  Sensor threshold crossed
  Button pressed
  Control loop timer expired
  Watchdog timed out

## Slide 22: Actions – Entry, Exit and Do
Entry action: executed immediately upon entering a state
Exit action: executed just before leaving a state
Do activity: ongoing activity during the state
Example:
Enter ‘infusing’ state: start motor
Exit ‘infusing’ state: stop motor
Promotes encapsulation: initialization, main logic and clean-up bundled together so easy to understand / debug / maintain
Prevents subtle bugs caused by repeated or missed actions

## Slide 23: Example state chart for the infusion pump
Making statecharts with Mermaid
Text-based way to create diagrams
Easy version control
Works directly in markdown files (on GitHub, VSCode with the ‘mermaid preview’ extension)

## Slide 24: The good and the bad state chart

## Slide 25: The good and the bad state chart

## Slide 26: What is wrong with the second example?
Do we need the extra states?
Every event is repeat for the ‘infusing’ and ‘paused’ state
Every new operational state multiplies transitions
Conceptual duplication: infusing / paused  alarm (over temperature) is conceptually the same thing, but not captured by the diagram
Safety is scattered (no single place that says “while infusing, safety monitoring is always active”)
Too complex looking
Is ‘paused’ a fundamentally different state, or a refinement of ‘infusing’?

## Slide 27: Hierarchical states
State Explosion: A complex system can easily have 50+ states, making the diagram unreadable.
Solution: super-states and sub-states
  A system can be in an OPERATIONAL super-state, which contains sub-states like SENSING and PROCESSING
Inheritance: events attached to super-state automatically apply to all sub-states

## Slide 28: The hierarchical infusion pump statechart

## Slide 29: Concurrent / parallel / orthogonal states
An embedded system might be doing multiple functions in parallel (multiple subsystems)
  Infusion control
  Alarm monitoring
  Communication
Statecharts allow concurrent regions
How to show synchronized behavior (for example via shared variables or internal events)?

## Slide 30: Infusing and logging in the infusion pump
Should logging be a concurrent state inside infusion, or separate altogether?
Might want logging to change behavior in safestate, so use a different logging substate there
Are we implementing two tasks / parallel code?
Only behavioral independence, not implementation

## Slide 31: Errors and fault states
Error: recoverable abnormal condition; fault: unrecoverable / safety-critical condition
Happy path modeling is only 20% of the work, rest is fault management; explicitly model error detection and transition to safe states
Safe state: predefined state that minimizes risk
Fault states: explicit states used to handle specific failures (ex, sensor disconnected)
Recovery path: Logic required to move from a fault state to an operational state; must be explicitly modeled

## Slide 32: Infusion pump safe state, fault state and recovery
Safe state
  Motor stopped
  Alarms active
  No infusion delivered
Fault state example
  Event: temperature crosses threshold
  Transition: over-temperature state
  Entry action: stop pump, sound buzzer
  Recovery: transition back to ‘infusing’ or ‘idle’ (design choice) when temperature back to normal

## Slide 33: Infusion pump statechart